## Slide 1: BACTEC MGIT 960 DST troubleshooting
<number>

## Slide 2: The right equipment used correctly
<number>
Start with the right equipment/supplies
  Use only sterile tubes and pipettes
  Sterile glass beads to break up clumps in inoculum from solid media
  Calibrated pipettes
    Drug addition to tubes
  0.5 McFarland standard with same tube diameter as inoculum tube and same media
    Organism in 7H9, use 7H9 blank in densitometer
  Vortex well and allow clumps to settle
    Do not select inoculum from bottom of tube or over inoculate tubes with clumps of organisms

## Slide 3: Use the correct reagents
<number>
PZA test medium and MGIT 7 ml tubes are not interchangeable when testing SIRE and PZA DST
PZA supplement and SIRE supplement are not interchangeable
Do not use supplement other than that supplied with the drugs – no PANTA or other growth supplement
Drug reconstitution:
  Use sterile distilled/deionized water
  Ensure drug dissolves completely
  Store reconstituted drug at -20 oC up to 6 months, but do not exceed original expiration date
  Use thawed, reconstituted drug the same day; discard any unused portions

## Slide 4: Use the correct inoculum
<number>
Start with the right inoculum
  Pure cultures only
  Fresh cultures
    Slant < 14 days
    Positive MGIT tube
      Day MGIT 960 called ‘positive’ plus 1 day to 5 days from date of positivity
  Homogeneous inoculum
    Use sterile glass beads and vortex to break up organism clumps
    Use only supernatant for inoculation
  Too high or too low an inoculum may give erroneous results or un-interpretable results
    Prepare dilutions according to procedure
    Use accurate pipette to add inoculum to tubes

## Slide 5: Error report for DST set
Some conditions occur that may affect the test results; the instrument will report this as an error (X)
  GC tubes become positive in less than 4 days
    Inoculum too heavy
  GC tubes remain negative up to 21 days
    Inoculum too light
<number>

## Slide 6: Invalid DST sets ‘X’
Growth control is not positive within 13 days (SIRE) or 21 days (PZA):
    Assumption is made that set is under-inoculated
Growth control is positive too soon (< 4 days):
    Assumption is made that set is over-inoculated or contaminated
ANY tube is missing in carrier
<number>

## Slide 7: DST interpretation of DST results
<number>
Invalid results ‘X’
  Initial inoculum was too heavy
    If working from MGIT tube, ‘Day’ 3 to 5 must be diluted
    If working from solid media
      Tube must be allowed to settle – over-inoculation and minimize aerosols
      Adjust suspension from solid to 0.5 McFarland
    Initial specimen not vortex-mixed completely; clumps still present in suspension
  Growth Control suspensions for SIRE and PZA are different (1:100 vs 1:10)

## Slide 8: DST troubleshooting: Non-interpreted results
<number>
Instrument took too long to interpret results (SIRE >14 days; PZA >21 days)
  Initial inoculum too light
    Precise pipettes must be used for dilutions and addition to tubes
      Not transfer pipettes
    If working with solid media
      Media may have been picked up, which interferes with reading 0.5 McFarland
  Initial culture too old
    Not in proper growth phase

## Slide 9: DST troubleshooting
<number>
False susceptible
  Organism inoculum is too light
    Starting suspension must be > 1.0 McFarland
  Suspension not homogeneous
  Culture too old
  Growth control dilution was too heavy
    Growth control grew too fast
  Was proper amount of supplement added to tubes?
    Was supplement properly stored?
  Drug not reconstituted properly
  Too much drug added to tubes
  Are tube caps on tight? If not, oxygen can leak into tube
  Only PZA medium and PZA supplement can be used for PZA test
  Organism could be borderline-resistant

## Slide 10: DST troubleshooting
<number>
False resistant
  Culture is contaminated or mixed culture
  Organism inoculum was too heavy
  Organism was not properly diluted
  Suspension was not homogeneous or had clumps
  Growth control dilution was too light
    Growth control grew too slowly
  Drug was not reconstituted properly
  Drug was not stored properly
  Drug was not added to tube
  MGIT tube was not mixed properly after the addition of the organism suspension
  Organism could be borderline-resistant

## Slide 11: DST troubleshooting
<number>

## Slide 12: Perform definitive identification
Perform ZN smear on positive MGIT tubes
Identify MTBC
  Some NTMs show resistance to anti-mycobacterial drugs
<number>

## Slide 13: Verify purity of inoculum
Confirm inoculum purity
  ZN smear
    Record cellular morphology on worksheet
      MTBC mixed with NTM
    Contamination (fungus and bacteria)
  Blood plate for contamination check
  Middlebrook 7H10
    Rule out mixed mycobacterial cultures
<number>

## Slide 14: Verify resistance profile
Any resistance to a drug
  Confirm resistance
    Drug not added to tube
    Less than 0.1 ml of drug added to tube
    Inoculum selected from the bottom of seed tube or clumps were not allowed to settle
  Inoculate a MGIT 960 seed tube
  Reset DST panel with resistant drug
<number>

## Slide 15: ZN on growing drug-containing tubes
If Middlebrook 7H10 is not available to confirm mycobacteria purity
  ZN on resistant drug-containing tubes
    Verify that only one type of cellular morphology is present
    Still need to reset DST to verify drug was added to tube
<number>

## Slide 16: DST troubleshooting: When to re-test
<number>
Un-interpreted results
Mono-resistance (especially RIF, EMB & PZA)
Results disagree with previous testing from the same patient
  Borderline resistance/susceptibility?
Patient is a new untreated case and shows resistance
QC organisms fail to give expected results
Unexpected high number of resistant isolates in a batch

## Slide 17: Troubleshooting summary
<number>
Organism suspension must be homogeneous, without clumps
Care must be taken when preparing dilutions
Drugs must be reconstituted and stored properly
Sterile tubes and pipettes must be used
Accurate pipettes must be used, not disposable transfer pipettes
QC of DST is critical to ensure test is functioning properly
  Should be performed on all new lots of DST reagents and MGIT tubes and weekly with each batch of testing

## Slide 18: Acknowledgement
<number>