## Slide 1: BACTEC™ MGIT™ 960 culture positive, AFB smear, and identification
<number>
Notes: <number>

## Slide 2: Training outline
<number>
Overview of BD MGIT 960 Instrument
Removing positive cultures from the instrument
Printing MGIT 960 unloaded positives report
Preparing and reading ZN AFB smear
Identification
Notes: <number>

## Slide 3: MGIT 960 drawers
Drawers designated as A, B & C
Sample Measurement Module (SMM)
  Tube Rack – “stations”
  Detector Assembly – 16 detectors for each row. Moves Left to Right.
  Status Indicators – three lamps on drawer
  Station Status LED – indicators on each station. Colors: red, green or orange

## Slide 4: MGIT 960 barcode scanner
Located on the front of the instrument to provide ability to scan tube labels for specimen identification.
Turns on automatically whenever the system is ready (and expecting) to scan a barcode

## Slide 5: LCD display and keypad
Presents information about system status and function key definitions.
Each display shows icons representing the current soft key assignments at the bottom of the screen

## Slide 6: BBL MGIT tube (7 ml)
<number>
Modified Middlebrook 7H9 broth base
Fluorescent indicator detects O2 level
  Embedded in silicon in bottom of tube
Bacteria use O2 during growth
  O2 released from indicator in tube bottom
  Fluorescence indicates growth of bacteria
Notes: <number>

## Slide 7: Indicator system
F
F
FO2
FO2
FO2
FO2
FO2
F
FO2
FO2
CO2
O2
O2
O2
O2
O2
O2
O2
O2
CO2
CO2
O2
F
F
F
F
FO2
FO2
F
F
F
F
CO2
O2
O2
O2
O2
O2
headspace
broth
sensor
Little or no fluorescence
Strong fluorescence
Negative culture
Positive culture
Notes: <number>

## Slide 8: Notification of positives on drawer
<number>
Indicator lamp on drawer illuminates
Audible alert sounds
+
–
!
Notes: <number>

## Slide 9: BACTEC™ MGIT™ 960 Instrument Interface
A / D17
#1
09/19/96
00:59
37.2 ° C
233 117 319
5 0 0
7 10 0
0 0 0
0 0 0
74 192 0
1 1 1
0 0 0
0 0 0
Notes: <number>

## Slide 10: Notification of positives on LCD display
<number>
#1
09/19/96
00:59
37.2 ° C
A / D17
5 0 0
7 10 0
0 0 0
0 0 0
74 192 0
1 1 1
233 117 319
Notes: <number>

## Slide 11: Instrument positive cultures
<number>
Open the drawer on which the positive indicator light is illuminated
Press the “remove positive tubes” soft key
Positive station LEDs will light (green)
Remove the tube from the instrument
Scan the barcode
Visually check tube for flaky, clumped growth
Repeat until instrument beeps 3 times signaling all positives have been removed
Notes: <number>
55

## Slide 12: Print unloaded positive report
      From the status screen, press the “print reports” soft key
      Press the soft key that corresponds to the “Unloaded Positive Tubes”
      Ensure that all tubes listed on the report correspond to the tubes removed
      Write date positive on the MGIT tube with permanent marker
<number>

## Slide 13: Unloaded positives report
<number>

## Slide 14: Unloaded positives report
<number>
Time to detection (TTD) must be ≥ 4 days
If TTD is less than 4 days, it is likely to be contaminated or an NTM

## Slide 15: Unloaded positive report
<number>
The MGIT culture must be 1 - 5 days old to be used as a seed for DST inoculation. The day it is determined to be positive by the instrument is day 0.

## Slide 16: Identification of specimens
Printed reports list the MGIT tube barcode number
Need to link MGIT tube barcode with the laboratory specimen number
  Label tube with an additional barcode specifying the laboratory specimen number (if available)
    Instrument can read a second barcode
    Second barcode number on printed report
If no specimen barcode label, specimen number must be linked with MGIT tube barcode number
  Handwritten on report next to the tube barcode number
<number>

## Slide 17: Unloaded positive report
<number>
Accession number is laboratory specimen number
Sequence number is tube barcode number

## Slide 18: AFB positive?
Prepare a smear of the positive MGIT 960 tube
  Label a new clean glass slide with specimen ID #, source (MGIT) and date
  Work in a BSC, in an isolation room; wear appropriate PPE
    Potential high concentration of M. tb
<number>

## Slide 19: AFB smear preparation
  Do not shake or invert the tube
  Remove a very small amount of broth from the bottom of the tube
    Use a fine tip sterile transfer pipette
  Express one drop on the slide and spread it in a 1 x 2 cm oval
  Allow to dry INSIDE the BSC
  Add a drop of fixative to help adhere the positive MGIT broth to the slide
    Fixative is usually based on bovine or rabbit serum
<number>

## Slide 20: Heat fix and stain with Ziehl –Neelsen reagents
Heat-fix the smear using a hot plate (65–80 ˚C for 2 hours or with a flame (pass through 3–4 times)
ZN staining
  Carbol fuchsin - heated to steaming - 5 minutes
  3% Acid alcohol - 2 minutes
  Methylene blue - 1 minute
<number>

## Slide 21: Reading AFB smear
  Examine the stained smears with a bright field microscope using the 100x oil objective for the presence of acid-fast bacilli (AFB) in a methodical manner to ensure that an area is scanned only once
    scanning a series of three parallel sweeps the length of the smear, or by scanning a series of nine parallel sweeps the width of the smear
    Smears scanned in this manner will give about 60 to 100 consecutive fields in one sweep with a 100x objective
Looking for red bacilli, possible serpentine cords, on a blue background
<number>

## Slide 22: Definitive identification
An AFB + smear with serpentine cording is NOT a definitive identification system for reporting M. tuberculosis complex
  There are over 100 species of mycobacteria that are all AFB +
  M. kansasii also exhibits serpentine cords
<number>

## Slide 23: Definitive identification options
Chromatographic lateral flow immunoassay
    Capilia, SD Bioline, or MTBc
  Line Probe Assay
    CM or MTBDR+
  GeneXpert MTB/RIF
<number>

## Slide 24: Antigen detection using chromatographic immunoassay
Procedure:
Note: Unlike HIV rapid tests, this test must be inoculated, read and discarded inside a BSC!
A 100 µl suspension of growth from liquid or solid medium is added to the specimen well and, if belonging to the TB complex, will react with the Ab to MPB64
The test is read after 15 minutes

## Slide 25: Test procedure
Perform test in a BSC; wear appropriate PPE
Label test cartridge
Dispense 100 µl of the well-mixed MGIT culture with a transfer pipette on the specimen placing area of the cartridge
Examine the reading area of the cartridge after 15 minutes
Note: Test results must be interpreted no later than 60 minutes after inoculation of the test cartridge due to the inoculum drying, which may alter results
<number>

## Slide 26: Quality control
Done each day test is performed and documented on log sheet
Positive control
  Well-characterized MTBC isolate or a reference strain growing in liquid media
Negative control
  Well-characterized strain or reference strain of non-tuberculous mycobacteria (NTM)
    M. avium, M. fortuitum, etc. growing in liquid media
<number>

## Slide 27: Interpretation of Results
Positive
  formation of a purple to red line on the reading areas labeled [T] and [C] of the cartridge
Negative
  formation of a purple to red line on the reading area labeled [C] of the cartridge but not [T]
Invalid
  if no line is observed on the reading area [C], technical errors or product damage has occurred. In this case the test should be considered invalid and repeated using a new cartridge
<number>

## Slide 28: Antigen detection using the chromatographic immunoassay
Interpretation:
MTB complex – two purple-reddish bands at test and control sites
NTM or other bacterial species – one purple-reddish band only at control site
Note: The test cannot be interpreted if there is no band at the control site
<number>

## Slide 29: Positive capilia result (photo kindly provided by BMA laboratory, Bangkok)
<number>

## Slide 30: Summary
<number>
BACTEC MGIT 960 – decreased time to detection for growth
Prepare a smear to confirm MGIT tube growth is AFB positive and rule our contaminants
Identify all AFB growth as MTBC before setting up DST
  Rapid ID system – lateral flow
  Molecular identification

## Slide 31: Acknowledgement
<number>